## Introduction
INTRODUCTION
2025 CONTRACTOR LABORATORY QUALIFICATION MANUAL
3) When the lab is ready for on-site review, contact the local NDOT QA Manager or
This manual has been developed to outline the process for qualifying Flexible Pavements Quality |  | Jacob Reynolds in the Hot Mix Aspahlt Testing Laboratory (531) 249-8692 to schedule the
Control laboratories (permanent mix design labs and field laboratories) to perform testing for projects |  | on-site qualification. The completed manual including,
for the Nebraska Department of Transportation or have funding or oversite by the FHWA and to provide a |  | certifications, verifications, including those from an outside source must be fully completed and
a system to insure compliance with NDOT testing laboratory and acceptance standards. Any laboratory |  | electronically submitted by April 1st of each calendar year. All on-site reviews of the Contractor's Mix
not meeting the requirements set forth in this document is not qualified to perform laboratory testing |  | Design Laboratory may be performed up to 6 months prior to the full final submittal date. The April 1st
services on such projects. |  | requirement is in recognition that some of the annual calibration requirements are from outside entities,
and the on-site inspection is likely to precede the Gyratory, Scales, Thermometers etc.. certifications
The following manual consists of procedures for the calibration of laboratory equipment along with |  | and calibrations.
documentation for the associated laboratory equipment. The test technicians must be certified or
provisionally certified as required in NDOT's Materials Sampling Guide and Specifications. |  | 4) The on-site review will be performed at the Contractor's Mix Design Laboratory where a proficiency
sample will be tested. All the other field laboratories must have this Qualification documentation
The qualification procedures will require either a certificate, verification, or calibration for the |  | completed as well, these field laboratories will be reviewed when Independent Assurance reviews are
specified testing equipment, as defined below: |  | conducted.
Certification: A certification of calibration are those received with the purchase of the standard |  | 5) In the event that equipment or test do not meet certification, calibration, or proficiency test tolerance
test equipment that is used to calibrate or verify other equipment in the |  | and/or standards, the laboratory equipment shall be repaired, corrected, replaced, or taken out of service.
laboratory, or those received when an outside contractor calibrates testing |  | Documentation of the corrective action taken for the deficiencies shall be documented in this manual
equipment, or those received from the manufacturer when new equipment is |  | within 30 days of either notification and findings by NDOT, or findings by the contractor as part of their
purchased. |  | laboratory QC program.
Verification: A verification is performed by the laboratory and its technicians and is then |  | It is the Contractor's sole responsibility to maintain the equipment according to the procedures set forth
reviewed by the NDOT Laboratory Representatives annually. Complete calibration |  | in this manual and by any additional means that they deem necessary over and above these procedures
and verification records shall be kept on file in the laboratory at all times. |  | to provide and strengthen a robust QC program.
Calibration: These calibration and measurement records must be performed by the contractor |  | This manual is intended to define a process to verify and document that a laboratory is qualified and is
for review by NDOT Laboratory Representatives. |  | acceptable to perform testing for NDOT projects operating under the QA/QC testing program for NDOT
approved testing laboratories standards for FHWA, State and Federally Funded Projects for Asphaltic Concrete.
NOTE - All technician certifications, equipment certifications and calibrations from outside entities |  | This qualification is for the purposes stated previously and is provided (as is) with no warranty, expressed or
shall be scan/copied and placed in the attached file folder labeled "Certifications" with this |  | implied, including but not limited to: certifications, verifications, or procedures are disclaimed, and in no
Laboratory manual file folder. |  | event shall NDOT, FHWA, or any entities be liable to the contractor/laboratory or to any third party for any
direct, indirect, incidental,consequential,special, or exemplary damages or lost profit resulting from any use
The procedure for the qualification process is as follows: |  | or misuse of this data, incorrect data or any data produced as a product of this material. All procedures,
spreadsheets, and references are for informational and educational use only, and shall not be reproduced
1) The contractor/laboratory receives a copy of this manual prior to NDOT's review. |  | for any other purpose than its intended use.
2) The contractor/laboratory goes through each procedure for all of their equipment
and assembles the proper documentation as defined in each procedure in this manual.

## Evaluation Checklist
2025 LABORATORY EVALUATION CHECKLIST
CONTRACTOR
LABORATORY I.D.
Mailing Address
City, State, Zip
*Contact Person
E-MAIL Address
Date Laboratory Certified
PERSONNEL INFORMATION |  | PHONE NUMBERS
NAME | POSITION | LAB PHONE |  | FAX | CELL PHONE |  | E-MAIL ADDRESS
CERTIFIED LABORATORY TECHNICIANS
|  | ASPHALT | [5 empty] | ASPHALT
NAME |  | 5 (YR) | I.A. (1 YR) |  | NAME |  |  | 5 (YR) | I.A. (1 YR)
EQUIPMENT CALIBRATED BY -
TYPE OF LABORATORY ACCREDITATION | [5 empty] | REPORTS RECEIVED |  | DATE -
NDOT
REMARKS

## Lab Evaluation
2025 EVALUATION OF CONTRACTORS LABORATORY
NEBRASKA DEPARTMENT OF TRANSPORTATION
LABORATORY |  |  |  |  | Date Evaluated
Location of Lab. |  |  |  |  | Date Procedure Reviewed
Phone Number |  |  |  |  | Lab. Personnel
E-Mail Address
Fax Number
Inspected By
THIS IS TO BE COMPLETED BY NDOT LABORATORY REPRESENTATIVE
1. Condition of Laboratory | Very good |  | Good |  | Fair | Poor
2. Condition of Equipment | Very good |  | Good |  | Fair | Poor
3. Calibration I.D. Tags on Equipment |  | Yes |  | No
Remarks
4. Calibration Certificates Received |  | Yes |  | No
Remarks
5. Calibration of Equipment is Performed By.
Remarks
6. Certified to Test what Material
| Asphalt |  | Concrete |  | Soil | Aggregate
7. Received Technicians Testing Certificates |  |  | Yes |  | No
Site Manager
Remarks
8. Cooperation from Lab. Personnel
| Very good |  | Good |  | Fair | Poor
Remarks
9. Lab. Personnel Reviewed
(1 Year I.A.)
10. Lab. Personnel Reviewed
(5 Year Certificate)
11. Testing Procedures Reviewed |  | Complete Asphalt Sample |  |  | Yes | No
Remarks
12. Quality Assurance Manual Kept on File |  |  | Yes |  | No
Not Required
13. Laboratory AMRL Inspected | Yes |  | No
Remarks

## Nuclear Gauge or PQI
2025 DENSITY GAUGE EVALUATION CHECKLIST (NUCLEAR OR PQI)
RADIATION SAFETY
OFFICER (if applicable)
Mailing Address
City, State, Zip
Phone Number
E-MAIL Address
GAUGE INFORMATION |  | LAST DATE OF CALIBRATION | NEXT CALIBRATION DUE DATE | GAUGE STORAGE LOCATION
MODEL | SERIAL NUMBER
EQUIPMENT CALIBRATED BY -
CERTIFIED NUCLEAR GAUGE TECHNICIANS
| INITIAL SAFETY TRAINING DATE | RADIATION SAFETY CERTIFICATION AND EXP. DATE | DOT/HAZMAT CERTIFICATION AND EXP. DATE
NAME
REMARKS

## Procedure #35 FAA
Procedure for Verifying Funnel Stand Apparatus and
100 ml Cylindrical Measure
Procedure No. 35
Equipment Checked: |  | FUNNEL STAND APPARATUS AND 100 ml CYLINDRICAL
MEASURE (AASHTO T 304)
Purpose:
This procedure provides instructions for checking the volume of the cylindrical measure
and the critical dimensions of the funnel stand apparatus used in this test method.
Inspection Equipment Required:
1. Caliper
2. Balance
3. Thermometer
4. Glass plate and grease
Tolerances:
The equipment shall meet the tolerances specified in the test method
AASHTO T 304.
Procedure:
The first step is to calibrate the cylindrical measure. Apply a light coat of grease
to the top edge of the dry empty cylinder measure. Weight the measure, grease
and glass plate. Fill the measure with deionized water at the temperature of 18
to 24° C (64.4 to 75.2° F). Record the temperature of the water. Place the glass
plate on the measure, being sure no air bubbles remain. Dry the outside of the
measure and weigh the measure with the water and glass plate on top. Calculate
the volume of the measure on the worksheet.
After calibrating the measure, place the measure on the funnel stand apparatus
and check the critical dimensions. Check the dimensions with the tolerances
specified in AASHTO T 304 test procedure.
Verification Interval:
12 Months
Report:
A copy must be retained in the Testing Facilities files for review.

## #35 Worksheet FAA
Uncompacted Void Content of Fine Aggregate
Procedure No. 35
AASHTO T 304
Date: |  |  |  | Location of Equipment
Checked by: | [5 empty] | Equipment Serial No.:
CALIBRATING FAA CUP
TEMPERATURE CHART
Wt of Cup-Glass-Grease with Water | [6 empty] | C | F | kg/m3 | lb/ft3
[7 empty] | 15.6 | 60 | 999.01 | 62.366
Wt of Cup-Glass-Grease | [6 empty] | 18.3 | 65 | 998.54 | 62.336
[7 empty] | 21.1 | 70 | 997.97 | 62.301
Net Mass of Water, (gm.) |  |  |  | (M) |  | (M) | 23 | 73.4 | 997.54 | 62.274
[7 empty] | 23.9 | 75 | 997.32 | 62.261
Density of H2O kg/m3 (Take off Chart) |  |  |  | (D) |  | (D) | 26.7 | 80 | 996.59 | 62.216
[7 empty] | 29.4 | 85 | 995.83 | 62.166
Volume of FAA Cup (nearest 0.1) |  |  |  | mL (V) |  | mL (V)
[6 empty] | 1 | 2
Calculation | V=1000 | M |  |  | Avg. of 2
|  | D |  |  |  |  | Avg of 2
Apply a light coat of grease to top edge of the dry, empty cylindrical measure. Weigh the
measure,grease, and glass plate. Fill the measure with freshly boiled, deionized water at a
temperature of 18 to24 degree C. Record the temperature of the water.Place the glass plate on the
measure, being sure no air bubbles remain. Dry the outer surface of the measure and determine
the combined mass of measure, glass plate, grease, and water by weighing. Following the final
weighing, remove the grease, and determine the mass of the clean,dry,empty measure
for subsequent tests. |  |  | AASHTO T 304-5 (8)
Cup Number
Weight of Cup
Volume of Cup
CYLINDER DIMENSIONS
Outside diameter | [7 empty] | mm (41 mm approx.)
Inside diameter | [7 empty] | mm (39 mm approx.)
Inside height | [7 empty] | mm (86 mm approx.)
Outside height | [7 empty] | mm
Bottom of Cylinder | [7 empty] | mm (Min. 6mm Thick)
Funnel Stand Apparatus w/ Cylindrical Measure in Place
Funnel height |  |  | mm (Min. 38 mm)
Funnel opening |  |  | mm (12.7 + 0.6 mm dia.)
From funnel opening to top of cylindrical measure | [5 empty] | mm (115 + 2 mm)

## Procedure #60 Splitter
Procedure for Verifying Mechanical Splitters
Procedure No. 60
Equipment Checked: |  | MECHANICAL SPLITTERS (AASHTO R 76)
Purpose:
This method provides instructions for checking mechanical sample splitters used in
AASHTO Test Method R 76.
Inspection Equipment Required:
1. Ruler.
Tolerance:
Sample splitters shall have an even number of equal width chutes, but not less than
a total number of eight and twelve chutes for coarse and fine aggregate, respectively.
The pouring pan shall have a width equal to or slightly less than the overall width of the
assembly of the chutes. The Receptacle pans shall have a width slightly greater than the
total width of the splitter chutes.
Procedure:
1. Measure the length of the divider with a ruler and record length and condition.
2. Measure the length of the pouring pan with a ruler and record length and condition.
3. Measure the length of the Receptacle pans with a ruler and record length and condition.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #60 Worksheet Splitter
Sample Splitter
Procedure No. 60
AASHTO R 76
Date : |  |  | LOCATION OF EQUIPMENT
Checked by:
Calibration Equipment Used:
Splitter - #1
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Splitter - #2
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Splitter - #3
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Remarks:

## Procedure #38 Ignition
Procedure for Verifying Ignition Oven
Procedure No. 38
Equipment Checked: |  | IGNITION OVEN (AASHTO T 308)
Purpose:
This procedure provides instructions for calibrating the ignition oven balance.
Inspection Equipment Required:
1. 8000 gram calibration mass.
Tolerance:
The equipment checked shall meet the tolerances specified in the test method listed
above.
Procedure:
1. For calibrating the ignition oven balance, follow the Preventive Maintenance and
Servicing Manual procedure.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #38 Worksheet Ignition
IGNITION OVEN
CERTIFICATE OF VERIFICATION
Procedure No. 38
AASHTO T 308
DATE -
EQUIPMENT VERIFIED -
EQUIPMENT SERIAL NUMBER -
METHOD USED FOR ACCEPTANCE - |  |  | Refer to Operator's Manual
(To Span Calibrate the Balance)
CALIBRATION EQUIPMENT USED -
VERIFICATION WEIGHTS - |  |  |  | 1000.0 gm
| (Read-out on Ignition Oven) |  |  | 3000.0 gm
5000.0 gm
8000.0 gm
EQUIPMENT CHECKED BY -
LOCATION OF EQUIPMENT -
This certificate is issued for the aforementioned equipment and certifies that the equipment has
been checked and found to comply with the requirements set forth in the method listed above.
In order for this equipment to remain in service it must be recertified on the following date and
must comply with the latest methods listed at that time.
NEXT CALIBRATION DUE -
Remarks

## Procedure #37 Gyratory
Procedure for Verifying Gyratory Compactors
Procedure No. 37
Equipment Checked |  |  | GYRATORY COMPACTOR ( AASHTO T 312)
Purpose:
This procedure provides for calibrating the gyratory ram pressure, internal angle of
gyration, gyration frequency and LVDT.
Inspection Equipment Required:
1. Calibration kit, including equipment to verify and calibrate the pressure and angle.
Tolerance:
The equipment checked shall meet the tolerance specified in the equipment owners
manual.
Procedure:
The procedure for calibrating the gyratory compactor is located in the Manual of
Operation and Maintenance for Troxler Model 4140 and 4141. All other gyratories
are calibrated by an outside source and will require certification documents from
those companies.
Verification Interval:
12 months
Report:
A copy of the certificate of calibration must be retained in the Testing Facilities files
for review.

## #37 Worksheet Gyratory
Gyratory Compactor
Procedure No. 37
AASHTO T 312
GYRATORY COMPACTOR CALIBRATION
LABORATORY NAME
LOCATION OF LABORATORY
TYPE OF GYRATORY | [5 empty] | MODEL OF GYRATORY
SERIAL NUMBER | [5 empty] | CALIBRATION KIT USED
CALIBRATION PERFORMED BY | [5 empty] | CALIBRATION KIT LAST CALIBRATED
DAVE-2 CALIBRATION KIT LAST CALIBRATED
COMPANY I.D. TAG # | [5 empty] | COMPACTOR LAST CALIBRATED
1 | PRESSURE CALIBRATION
Date
System Load Cell |  |  |  | kPa |  | TOLERANCE -- |  | 3 Percent or +/- 18 kPa
Calibration Load Cell |  |  |  | kPa
2 | ANGLE CALIBRATION |  |  |  | MOLDS USED
|  | Date |  |  |  |  | TOP | BOTTOM | Avg. | Tolerance
[5 empty] | Internal Angle |  |  |  | 0 | 1.16 +/- .02
Measured Internal Angle |  |  |  | Deg.
[6 empty] | External Angle (If Measured) |  |  |  | Deg.
Ram Head Measurement |  |  | A Reading | B Reading | C Reading |  |  |  |  | 1.25 +/- .02
[6 empty] | Mold Size |  |  | 150 | mm
3 | HEIGHT CALIBRATION
|  | Date |  |  |  | Daily File Kept |  |  |  | y/n
Actual |  |  |  | mm |  | Measured |  |  |  | mm
Pressure |  |  |  | kPa |  | Mold Size |  |  | 150 | mm
4 | ROTATION CALIBRATION |  |  |  | Performed as manufacturers recommendation.
Date
Target |  |  | 30 | RPM |  | Measured |  |  |  | RPM
5 | MAINTENANCE
Follow manufacturers recommended schedule of maintenance. (Includes vacuuming, degreasing, lubricating,
and inspecting for loose or worn parts)
Kept Clean |  |  |  | y/n |  | Well Greased |  |  |  | y/n
OTHER MAINTENANCE PERFORMED -
REMARKS

## Procedure #36 Molds
Procedure for Verifying Gyratory Molds
Procedure No. 36
Equipment Checked: |  | GYRATORY MOLDS AASHTO T 312 -Annex A
Purpose:
This procedure provides instructions for checking the critical dimensions of the molds
and the end plates used in this test method.
Inspection Equipment Required:
1. Calipers or steel rule capable of measuring to 0.01 millimeters.
2. Surface Comparator.
3. Rockwell Hardness Tester.
4. Three Point Bore Gauge or Coordinate Measurement Machine (CMM).
Procedure:
1. Measure and record the inside diameter of the mold to the nearest 0.001 mm. Rotate
the mold in 90 degree intervals, measure and record the inside diameter to get a minimum of
three measurements. Perfom this procedure at the top, bottom and maximum wear position.
2. Measure and record the length of the mold to the nearest millimeter. Repeat this
procedure at intervals of 90 degrees around the mold.
3. Measure the diameter of the mold end plate at intervals of 90 degrees around the plate.
4. Check the smoothness of the molds and bottom plate surfaces that are in contact with
the specimen. Use a surface comparator. (Follow owners manual)
5. Check the Rockwell Hardness of the molds and bottom plates with a Rockwell Hardness
tester. (Follow owners manual)
Note: Only new molds need the hardness tested.
6. Measure and record the diameter gyratory ram at intervals of 90 degrees to the nearest
millimeter.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #36 Worksheet Molds
Gyratory Molds
Procedure No. 36
AASHTO T-312
GYRATORY MOLD CALIBRATION
DATE ----
LOCATION OF EQUIPMENT --
EQUIPMENT CHECKED BY --
METHOD USED FOR ACCEPTANCE -- T-312 Annex A
CALIBRATION EQUIPMENT USED --
New molds inside diameter 149.90 to 150.00 mm. In-service molds inside diameter shall not exceed 150.2 mm.
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
Mold #
SURFACE FINISH OF MOLDS-- (1.60 um or smoother)
ROCKWELL HARDNESS (min. C48, One time only)
NEXT CALIBRATION DATE -

## Procedure #23 Gmm
Procedure for Verifying Vacuum System Pressure
Procedure No. 23
Equipment Checked: |  | VACUUM SYSTEM (AASHTO T 209)
Purpose:
This method provides instructions for checking the pressure developed by vacuum
pumps for the Rice Test.
Procedure:
1. Connect the vacuum pump to the vacuum chamber.
2. Place a manometer in series between the pump and the vacuum chamber being sure
all connections are tight.
3. Start the pump and allow it to run for at least 5 minutes to allow the system to stabilize.
4. Record the vacuum attained and the serial numbers of pump and manometer being used.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #23 Worksheet Gmm
VERIFYING VACUUM SYSTEM PRESSURE
Procedure No. 23
AASHTO T 209
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
VACUUM SYSTEMS 1-3
System No. | 1 |  |  | 2 |  |  | 3
Manometer Type
Pump Type
No. of Moisture Traps
Micro Valve Installed |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Bowl Standardized |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Wire Mesh on Lid Opening |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Water Bath |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N
Mechanical Shaker |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Scale Used (Type)
Vacuum Pressure Obtained (25.0-30.0mm Hg.)
|  |  | mm Hg |  |  | mm Hg |  |  | mm Hg
VACUUM SYSTEMS 4-6
System No. | 4 |  |  | 5 |  |  | 6
Manometer Type
Pump Type
No. of Moisture Traps
Micro Valve Installed |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Bowl Standardized |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Wire Mesh on Lid Opening |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Water Bath |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N
Mechanical Shaker |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Scale Used (Type)
Vacuum Pressure Obtained (25.0-30.0mm Hg.)
|  |  | mm Hg |  |  | mm Hg |  |  | mm Hg
Remarks:

## Procedure #8 Scales
Procedure for Verifying Scales, Balances and Balance Weights
Procedure No. 8
Equipment Checked: |  | SCALES AND BALANCES (AASHTO M231)
Purpose:
This method provides instructions for checking the accuracy sensitivity of scales, balances
balance weights used in all labs.
Inspection Equipment Required:
1. Calibrated weights with certificates.
2. Powder free gloves designed for handling these weights.
Tolerance:
All AASHTO and ASTM tests requiring the use of scales and or balances have in their
procedures a list of requirements for the weighing devices. It will be up to the inspector
to check these requirements when checking and verifying this equipment.
Procedure:
Each balance or scale will have its own worksheet. Choose the appropriate worksheet for
the balance or scale to be verified. Check the equipment through the entire range on the
worksheet. It is recommended to calibrate the balance before verifying. Follow the owners
manual for this procedure.
Lab balance weights used only for verification purposes shall be verified against calibrated
weights.
Verification Interval:
12 months
Report:
A copy for each scale or balance must be retained in the Testing Facilities files for review.

## #8 Balance Weights Worksheet
Balance Weights
Procedure No. 8
AASHTO M 231 Table 4
Date :
Serial No.:
Checked by:
Material:
Scale No.:
Calibration Equipment:
Balance Weights
Weight | Accuracy
5000g | .80g
2000g | .40g
1000g | .25g
500g | .20g
200g | .10g
100g | .07g
50g | .04g
20g | .02g
10g | .015g
5g | .010g
2g | .006g
1g | .004g
Remarks:

## #8 Gen. Purp. Worksheet
General Purpose Scales
Procedure No. 8
Class G2
Date:
Location of Scale:
Inspected By:
Calibration Weights:
Scale No.:
Manufacturer:
Scale Capacity:
Test Point |  | Accuracy Required | Accuracy
50 g |  | .2 g
100 g |  | .2 g
200 g |  | .2 g
300 g |  | .3 g
400 g |  | .4 g
500 g |  | .5 g
1000 g |  | 1.0 g
2000 g |  | 2.0 g
3000 g |  | 3.0 g
4000 g |  | 4.0 g
5000 g |  | 5.0 g
6000 g |  | 6.0 g
7000 g |  | 7.0 g
8000 g |  | 8.0 g
9000 g |  | 9.0 g
10000 g |  | 10.0 g
11000 g |  | 11.0 g
12000 g |  | 12.0 g
13000 g |  | 13.0 g
14000 g |  | 14.0 g
15000 g |  | 15.0 g
16000 g |  | 16.0 g
Remarks:

## Procedure #1 Ovens
Procedure for Verifying Ovens
Procedure No. 1
Equipment Checked: |  | DRYING OVENS
Purpose:
This method provides instructions for checking drying ovens used in the laboratory.
Inspection Equipment Required:
1. A calibrated thermometer either Fahrenheit or Celsius graduated in 1.0° increments
having a range which includes the temperature range to be checked.
2. A thermometer well to retain heat while the oven door is open.
3. A clothes pin the hold the thermometer in such a manner as to enable the operator
to read the scale easily.
Tolerance:
Drying ovens shall be capable of maintaining a constant temperature range listed in the
appropriate test methods.
Procedure:
1. Place the thermometer inside the well with the clothes pin attached to the
thermometer. Position the thermometer on the shelf where the samples are normally
dried.
2. Take the first reading at least 1 hour after closing the oven (oven should remain
undisturbed).
3. Take as many readings as necessary to determine if the temperature range is within
the specified tolerance (three consecutive readings, taken no less that 1/2 hour apart,
within the tolerance allowed are adequate).
4. Adjust the temperature of the oven if an observed reading is outside the tolerance
specified (allow at least 1/2 hr. for the temperature to stabilize between each
adjustment. Return to step 3.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #1 Worksheets Ovens
Ovens
Procedure No. 1
AASHTO T-209, T-216, T-265, T-296
Date - |  |  | Inspected by:
Location of Equipment -
Calibration Equipment Used -
Oven #1 |  |  |  | Oven #2
Manufacturer |  |  |  | Manufacturer
Model No. |  |  |  | Model No.
Serial No. |  |  |  | Serial No.
Oven Set Point |  | Temperature Reading |  | Oven Set Point |  | Temperature Reading
| F | 30 min. |  |  | F | 30 min.
|  | 60 min. |  |  |  | 60 min.
|  | 90 min. |  |  |  | 90 min.
( 230+/- 9 Degrees F or 110 +/- 5 Degrees C) |  |  |  | ( 230+/- 9 Degrees F or 110 +/- 5 Degrees C)
Remarks |  |  |  | Remarks
Oven #3 |  |  |  | Oven #4
Manufacturer |  |  |  | Manufacturer
Model No. |  |  |  | Model No.
Serial No. |  |  |  | Serial No.
Oven Set Point |  | Temperature Reading |  | Oven Set Point |  | Temperature Reading
| F | 30 min. |  |  | F | 30 min.
|  | 60 min. |  |  |  | 60 min.
|  | 90 min. |  |  |  | 90 min.
( 230+/- 9 Degrees F or 110 +/- 5 Degrees C) |  |  |  | ( 230+/- 9 Degrees F or 110 +/- 5 Degrees C)
Remarks |  |  |  | Remarks
Remarks

## Procedure #27 Shaker
Procedure for Verifying Mechanical Sieve Shakers
Procedure No. 27
Equipment Checked: |  | MECHCANICAL SIEVE SHAKER (AASHTO T 27)
Purpose:
This method provides instructions for checking the sieving efficiency of mechanical
sieve shakers.
Inspection Equipment Required:
1. A scale capable of weighing to the nearest 0.1 gram.
2. An oven for drying the sample to a constant weight.
3. Sieves of sufficient diameter and screen openings so as not to cause overloading of
any one screen.
4. Sample splitter.
Procedure:
1. Dry the sample to a constant weight.
2. Split the dried sample so as to provide at least 500 grams of fine aggregate or 2500 grams
of coarse aggregate.
3. Place the prepared and weighed sample in the sieve nest and place on the shaker.
4. Shake the sample for 5 minutes and remove.
5. Hand shake each sieve over a pan for one minute and weigh the material passing and
record the weight for each sieve,
6. The material passing from hand shaking in Step 5 shall not exceed 0.5% of the total
weight of the sample used.
7. If the weight exceeds the requirement in #6, increase the time on the shaker and repeat
the test until the requirement is met.
8. When checking shaker which have adjustable shaking speeds it may be necessary to
adjust the speed as well as the time in order to meet the requirements.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #27 Worksheet Shaker
Mechanical Sieve Shakers
Sieving Thoroughness
Procedure No. 27
AASHTO T 27
Date:
Checked by:
Shaker Serial No.:
Manufacturer:
Type of Aggregate used: |  |  | Fine |  | Coarse
Scale Serial Number:
Oven Serial Number:
Sieve Diameter: |  | 8" |  | 10" |  | 12"
Initial Sample Weight: |  |  |  | gm.
Sieve |  | Screen |  | Material |  |  | Percent
Serial No. |  | No. |  | Passing |  |  | Passing
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
Total Mechanical Shaker Time |  |  |  | min.
NOTE: Use a new worksheet for each run and or type of aggregate used.

## Procedure #25 Sieves
Procedure for Verifying Sieves
Procedure No. 25
Equipment Checked: |  | SIEVES (AASHTO ASTM E-11)
Purpose:
This method provides instructions for visually checking sieves used in the laboratory.
Inspection Equipment Required:
1. A method for providing a lighted background when viewing the sieve screen.
2. A magnifier for close examination of the sieve screen.
Tolerance:
The sieves examined shall meet the requirements set forth in ASTM E-11.
Procedure:
1. Check each sieve tag for the proper designation and standard opening size. This
designation shall be in metric dimensions and in the case of older sieves the English
and metric dimensions shall be both on the tag.
2. Check the frame for excessive wear along the rim and for looseness between the top
and bottom halves.
3. Check the screen for loose wires, bowed screens, torn screens, creases I n the screen,
broken screen wires, and deformation of the screen openings. These items are normally
visible with the naked eye when viewed against a lighted background. The use of a
magnifier will aid when viewing the finer screens.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #25 Worksheet Sieves
Wire Cloth Sieves for Testing Purposes
Procedure No. 25
ASTM E-11
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Sieves
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Remarks:

## Procedure #64 Thermometer
Procedure for Verifying Thermometers
Procedure No. 64
Equipment Checked: |  | Thermometers
Purpose:
This method provides instructions for checking thermometers used in the laboratory and in
the field.
Inspection Equipment Required:
1. A calibrated thermometer either Fahrenheit or Celsius graduated in 1.0° increments
having a range which includes the temperature range to be checked. A calibrated
digital temperature measuring device is also acceptable.
Tolerance:
The tolerance for mercury thermometers is plus or minus one scale division of that
thermometer. Tolerance for spirit filled thermometers is plus or minus 1% of the total range
for that thermometer. Tolerance for dial thermometers is plus or minus one scale division.
Digital thermometers such as Oaktrons & Taylors used most commonly in water baths and in
ovens have a tolerance of plus or minus 3.6 degrees.
Procedure:
Verify lab or field thermometers against a calibrated thermometer. Check each thermometer
at two points if possible. Using water to perform the checks would be the best method.
It does no have to be a commercially made water bath. It could be as simple as using a
a bucket and adding hot water. A piece of Plexiglas with drilled holes works well when
placed over a water bath. Clothes pins or other clips can be used to hold the thermometers
in place while performing the check. Alternative methods may be used to achieve results.
Verification Intervals:
12 months
Report:
A copy must be retained for each thermometer tested and kept in the Testing Facilities
files for review.

## #64 Worksheet Thermometer
Thermometer
Procedure No. 64
ASTM E 77-07
Date:
Checked By:
Calibration Equipment Used:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks: