## Introduction
INTRODUCTION
2025 CONSULTANT LABORATORY QUALIFICATION MANUAL |  |  |  | Nebraska Department of Transportation
Robert C. Rea, Assistant Materials and Research Engineer
This manual has been developed to outline the process for qualifying a laboratory to perform |  |  |  | 1400 Highway 2
testing for projects for the Nebraska Department of Transportation or have funding or oversite by the |  |  |  | P. O. Box 94759
FHWA and to provide a system to insure compliance with NDOT testing laboratory standards. |  |  |  | Lincoln, NE 68509-4759
Any laboratory not meeting the requirements set forth in this document is not qualified to |  |  |  | robert.rea@nebraska.gov
perform laboratory services on such projects. |  |  |  | Office (402) 479-3839
Cell (402) 430-4726
The following manual consists of procedures for the calibration of laboratory equipment along |  |  |  | Fax (402) 479-3975
with calibration worksheets (where required) as pertaining to the following disciplines:
Aggregate, Bituminous, PCC, Soils, and the associated laboratory equipment. The testers must |  |  | 4) The qualification review will take approximately 2 days, in which the laboratory must
be certified and perform tests as outlined in NDOT's QAP manual. |  |  | have their Laboratory Contact working directly with the NDOT lab qualification review
personnel throughout the review.
The qualification procedures will require either a certificate, verification, or calibration for the
specified testing equipment. |  |  | Certificates of calibration and the verification documentation must be performed at the
intervals specified by the procedures. Completed calibration and verification records shall be
Certificate: A certificate of calibration are those received with the purchase of the standard |  |  | kept on file for review at any time.
test equipment that is used to calibrate or verify other equipment in the
laboratory, or those received when an outside contractor calibrates testing |  |  | If the laboratory equipment does not meet the qualification requirements, the NDOT
equipment, or those received from the manufacturer when new equipment is |  |  | Laboratory Representatives shall notify the Laboratory Contact. After notification, the
purchased. |  |  | laboratory equipment shall be corrected, replaced, or taken out of service. Documentation of
the correction of the deficiencies shall be submitted to the NDOT Laboratory Representative
Verification: A verification is performed by the laboratory and its technicians and is then |  |  | within 30 days of notification, and the qualification process will resume.
reveiwed by the NDOT Laboratory Representatives at regular intervals as
identified in the procedures. Complete calibration and verification records shall |  |  | It is the laboratories sole responsibility to maintain the equipment according to the procedures
be kept on file at all times. |  |  | set forth in this manual and by any means they deem necessary over and above these
procedures.
Calibration: These calibration and measurement records must be performed by the consultant
for review by NDOT laboratory representative. |  |  | This procedure is for the sole purpose of defining a process to verify that a laboratory is
qualified and is acceptable to perform testing for NDOT Projects operating under the umbrella
of NDOT's approved testing laboratories standards for FHWA, State and Federally Funded
The procedure for the qualification process is as follows: |  |  | Projects. This qualification is for the purposes stated previously and is provided (as is) and any
express or implied warranties including, but not limited to, certifications, verifications, or
1) The consultant/laboratory receives copy of this manual prior to NDOT's review. |  |  | procedures are disclaimed. In no event shall NDOT, FHWA, or any entities be liable to the
consultant/laboratory or to any third party for any direct, indirect, incidental, consequential,
2) The consultant/laboratory goes through each procedure for each equipment and test |  |  | special, or exemplary damages or lost profit resulting from any use or misuse of this data,
and assembles the proper documentation as defined in each procedure in this manual. |  |  | incorrect data or any data produced as a product of this manual.
3) When all documentation is complete, contact the assistant Materials and Research
Engineer to schedule the qualification review at:
1 of 2 Pages

## Evaluation Checklist
2025 LABORATORY EVALUATION CHECKLIST
LABORATORY
Mailing Address
City, State, Zip
*Contact Person
E-MAIL Address
Materials Certified To Test |  | CONCRETE | ASPHALT |  | AGGREGATE |  | SOILS
PERSONNEL INFORMATION |  | PHONE NUMBERS
NAME | POSITION | LAB PHONE | FAX | CELL PHONE |  | E-MAIL ADDRESS
CERTIFIED LABORATORY TECHNICIANS
| SOILS | AGGREGATES | ASPHALT |  | CONCRETE
NAME |  |  | 5 (YR) | OTHER | ACI (5 YR) | I.A. (1 YR) | PLANT | OTHER
EQUIPMENT CALIBRATED BY -
TYPE OF LABORATORY ACCREDITATION |  |  |  |  | REPORTS RECEIVED |  | DATE -
REMARKS

## Nuclear Gauge
2025 NUCLEAR GAUGE EVALUATION CHECKLIST
RADIATION SAFETY
OFFICER
Mailing Address
City, State, Zip
Phone Number
E-MAIL Address
NUCLEAR GAUGE INFORMATION |  | LAST DATE OF CALIBRATION | NEXT CALIBRATION DUE DATE | NUCLEAR GAUGE STORAGE LOCATION
MODEL | SERIAL NUMBER
EQUIPMENT CALIBRATED BY -
CERTIFIED NUCLEAR GAUGE TECHNICIANS
| INITIAL SAFETY TRAINING DATE | RADIATION SAFETY CERTIFICATION AND EXP. DATE | DOT/HAZMAT CERTIFICATION AND EXP. DATE
NAME
REMARKS

## Equip. Cal. Info.
Equipment Calibration and Verification Information
LABORATORY:
LOCATION:
All calibration and verification records shall be kept in the office of the Commercial Laboratory Supervisor.
Equipment - Requirement - Test Method |  |  | Certificate | Procedure Reference
Verify
Calibrate
AGGREGATE EQUIPMENT
Stand, Funnel and Cylindrical Measure |  |  | Verify | Procedure No. 35
AASHTO T 304
Conical Molds and Tampers - Critical Dimensions |  |  | Verify | Procedure No. 9
AASHTO T 84
Sample Splitter |  |  | Verify | Procedure No. 60
AASHTO R 76
BITUMINOUS AGGREGATE EQUIPMENT
Ignition Oven |  |  | Verify | Procedure No. 38
AASHTO T 308
Gyratory Compactor |  |  | Certificate | Procedure No. 37
AASHTO T 312
Gyratory Molds and Ram Heads |  |  | Verify | Procedure No. 36
AASHTO T 312
Vacuum System |  |  | Calibrate | Procedure No. 23
AASHTO T 209
PORTLAND CEMENT CONCRETE EQUIPMENT
Air Meters |  |  | Verify | Procedure No. 61
ASTM C 231
Bearing Blocks and Retainers - Planeness |  |  | Verify | Procedure No. 11
ASTM C 39
Compression Machines |  |  | Certificate | Outside Contractor
ASTM C 39
Moist Rooms/ Storage Tanks - Temp. & Humidity |  |  | Verify | Procedure No. 62
ASTM C 39
Capping Material - Strength |  |  | Verify | Procedure No. 63
Unit Weight Measure |  |  | Verify | Procedure No. 26
ASTM C 138
Volumetric Air Meter |  |  | Verify | Procedure No. 65
ASTM C 173
Equipment - Requirement - Test Method |  |  | Certificate | Procedure Reference
Verify
Calibrate
SOILS EQUIPMENT
Grooving Tools - Critical Dimensions |  |  | Verify | Procedure No. 18
AASHTO T 89
Liquid Limit Devices - Wear and Critical Dimensions |  |  | Verify | Procedure No. 18
AASHTO T 89
Manual Hammers - Weight and Critical Dimensions |  |  | Verify | Procedure No. 5
AASHTO T 99 & T 180 |  |  |  | Procedure No. 5A
Mechanical Soil Rammers |  |  | Verify | Procedure No. 22
AASHTO T 99 & T 180 |  |  |  | Procedure No. 22A
Molds - Critical Dimensions/ Volume |  |  | Verify | Procedure No. 2
AASHTO T 99 & T 180
Straightedges - Planeness of Edge |  |  | Calibrate | Procedure No. 19
AASHTO T 99 & T 180
MISCELLANEOUS EQUIPMENT
Balances, Scales & Weights (General Purpose) |  |  | Verify | Procedure No. 8
AASHTO T 208, T 216, T 265, T 296
Ovens - Temperature Settings |  |  | Verify | Procedure No. 1
AASHTO T 208, T 216, T 265
Mechanical Shakers - Sieving Thoroughness |  |  | Verify | Procedure No. 27
Sieves - Physical Condition |  |  | Verify | Procedure No. 25
Thermometers |  |  | Verify | Procedure No. 64

## Lab Evaluation
2025 EVALUATION OF CONSULTANTS LABORATORY
NEBRASKA DEPARTMENT OF TRANSPORTATION
LABORATORY |  |  |  |  | Date Evaluated
Location of Lab. |  |  |  |  | Date Procedure Reviewed
Phone Number |  |  |  |  | Lab. Personnel
E-Mail Address
Fax Number
Inspected By
THIS TO BE COMPLETED BY NDOT LABORATORY REPRESENTATIVE
1. Condition of Laboratory | Very good |  | Good |  | Fair | Poor
2. Condition of Equipment | Very good |  | Good |  | Fair | Poor
3. Calibration I.D. Tags on Equipment |  | Yes |  | No
Remarks
4. Calibration Certificates Received |  | Yes |  | No
Remarks
5. Calibration of Equipment is Performed By.
Remarks
6. Certified to Test what Material
| Asphalt |  | Concrete |  | Soil | Aggregate
7. Received Technicians Testing Certificates |  |  | Yes |  | No
Site Manager
Remarks
8. Cooperation from Lab. Personnel
| Very good |  | Good |  | Fair | Poor
Remarks
9. Lab. Personnel Reviewed
(1 Year I.A.)
10. Lab. Personnel Reviewed
(5 Year Certificate)
11. Testing Procedures Reviewed |  | Complete Asphalt Sample |  |  | Yes | No
Remarks
12. Quality Assurance Manual Kept on File |  |  | Yes |  | No
Not Required
13. Laboratory AMRL Inspected | Yes |  | No
Remarks

## Procedure #35 FAA
Procedure for Verifying Funnel Stand Apparatus and
100 ml Cylindrical Measure
Procedure No. 35
Equipment Checked: |  | FUNNEL STAND APPARATUS AND 100 ml CYLINDRICAL
MEASURE (AASHTO T 304)
Purpose:
This procedure provides instructions for checking the volume of the cylindrical measure
and the critical dimensions of the funnel stand apparatus used in this test method.
Inspection Equipment Required:
1. Caliper
2. Balance
3. Thermometer
4. Glass plate and grease
Tolerances:
The equipment shall meet the tolerances specified in the test method
AASHTO T 304.
Procedure:
The first step is to calibrate the cylindrical measure. Apply a light coat of grease
to the top edge of the dry empty cylinder measure. Weight the measure, grease
and glass plate. Fill the measure with deionized water at the temperature of 18
to 24° C (64.4 to 75.2° F). Record the temperature of the water. Place the glass
plate on the measure, being sure no air bubbles remain. Dry the outside of the
measure and weigh the measure with the water and glass plate on top. Calculate
the volume of the measure on the worksheet.
After calibrating the measure, place the measure on the funnel stand apparatus
and check the critical dimensions. Check the dimensions with the tolerances
specified in AASHTO T 304 test procedure.
Verification Interval:
12 Months
Report:
A copy must be retained in the Testing Facilities files for review.

## #5 Worksheet FAA
Uncompacted Void Content of Fine Aggregate
Procedure No. 35
AASHTO T 304
Date: |  |  |  | Location of Equipment
Checked by: | [5 empty] | Equipment Serial No.:
CALIBRATING FAA CUP
TEMPERATURE CHART
Wt of Cup-Glass-Grease with Water | [6 empty] | C | F | kg/m3 | lb/ft3
[7 empty] | 15.6 | 60 | 999.01 | 62.366
Wt of Cup-Glass-Grease | [6 empty] | 18.3 | 65 | 998.54 | 62.336
[7 empty] | 21.1 | 70 | 997.97 | 62.301
Net Mass of Water, (gm.) |  |  |  | (M) |  | (M) | 23 | 73.4 | 997.54 | 62.274
[7 empty] | 23.9 | 75 | 997.32 | 62.261
Density of H2O kg/m3 (Take off Chart) |  |  |  | (D) |  | (D) | 26.7 | 80 | 996.59 | 62.216
[7 empty] | 29.4 | 85 | 995.83 | 62.166
Volume of FAA Cup (nearest 0.1) |  |  |  | mL (V) |  | mL (V)
[6 empty] | 1 | 2
Calculation | V=1000 | M |  |  | Avg. of 2
|  | D |  |  |  |  | Avg of 2
Apply a light coat of grease to top edge of the dry, empty cylindrical measure. Weigh the
measure,grease, and glass plate. Fill the measure with freshly boiled, deionized water at a
temperature of 18 to24 degree C. Record the temperature of the water.Place the glass plate on the
measure, being sure no air bubbles remain. Dry the outer surface of the measure and determine
the combined mass of measure, glass plate, grease, and water by weighing. Following the final
weighing, remove the grease, and determine the mass of the clean,dry,empty measure
for subsequent tests. |  |  | AASHTO T 304-5 (8)
Cup Number
Weight of Cup
Volume of Cup
CYLINDER DIMENSIONS
Outside diameter | [7 empty] | mm (41 mm approx.)
Inside diameter | [7 empty] | mm (39 mm approx.)
Inside height | [7 empty] | mm (86 mm approx.)
Outside height | [7 empty] | mm
Bottom of Cylinder | [7 empty] | mm (Min. 6mm Thick)
Funnel Stand Apparatus w/ Cylindrical Measure in Place
Funnel height |  |  | mm (Min. 38 mm)
Funnel opening |  |  | mm (12.7 + 0.6 mm dia.)
From funnel opening to top of cylindrical measure | [5 empty] | mm (115 + 2 mm)

## Procedure #9 Conical Mold
Procedure for Verifying Conical Molds and Tampers
Procedure No. 9
CONICAL MOLDS AND TAMPERS (AASHTO T 84)
Equipment Checked:
Purpose:
This method provides instruction for checking the critical dimensions of the conical mold
and the mass of the tamper used in this test procedure.
Inspection Equipment Required:
1. A scale or balance capable of weighing to 500 gram and readable to the nearest gram.
2. A caliper for measuring the inside diameter of the top and bottom of the cone and
measuring the outside diameter of the tamper face. The caliper should be able to read
to the nearest 0.1 mm.
3. A ruler for measuring the height of the cone and readable to the nearest 1.0 mm.
Procedure:
Using the requirements set forth in the AASHTO test method, check each of the
dimensions of the cone and the mass of the tamper.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #9 Worksheet Conical Mold
Specific Gravity and Absorption of Fine Aggregate
Procedure No. 9
AASHTO T 84
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Mold
Top inside diameter |  | Dia. # 1 |  |  | Dia. # 2 | (40 + 3 mm)
Bottom Inside diameter |  | Dia. # 1 |  |  | Dia. # 2 | (90 + 3 mm)
Height of mold |  | Ht. # 1 |  |  | Ht. # 2 | (75 + 3 mm)
Thickness of metal |  | Thick. # 1 |  |  | Thick. # 2 | (0.8 mm Min.)
Tamper
Weight of tamper |  |  | grams | (340 + 15 grams)
Diameter of the face |  |  | mm | (25 + 3 mm)
Is the face of the tamper flat? |  | Yes |  | No
Remarks:

## Procedure #60 Splitter
Procedure for Verifying Mechanical Splitters
Procedure No. 60
Equipment Checked: |  | MECHANICAL SPLITTERS (AASHTO R 76)
Purpose:
This method provides instructions for checking mechanical sample splitters used in
AASHTO Test Method R 76.
Inspection Equipment Required:
1. Ruler.
Tolerance:
Sample splitters shall have an even number of equal width chutes, but not less than
a total number of eight and twelve chutes for coarse and fine aggregate, respectively.
The pouring pan shall have a width equal to or slightly less than the overall width of the
assembly of the chutes. The Receptacle pans shall have a width slightly greater than the
total width of the splitter chutes.
Procedure:
1. Measure the length of the divider with a ruler and record length and condition.
2. Measure the length of the pouring pan with a ruler and record length and condition.
3. Measure the length of the Receptacle pans with a ruler and record length and condition.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #60 Worksheet Splitter
Sample Splitter
Procedure No. 60
AASHTO R 76
Date : |  |  | LOCATION OF EQUIPMENT
Checked by:
Calibration Equipment Used:
Splitter - #1
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Splitter - #2
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Splitter - #3
Splitter Brand -
Divider Width |  |  |  | Divider Condition
Pouring Pan
Width |  |  |  | Condition
Number of Pans Checked
Receptacle Pan
Width |  |  |  | Condition
Number of Pans Checked
Remarks:

## Procedure #38 Ignition
Procedure for Verifying Ignition Oven
Procedure No. 38
Equipment Checked: |  | IGNITION OVEN (AASHTO T 308)
Purpose:
This procedure provides instructions for calibrating the ignition oven balance.
Inspection Equipment Required:
1. 8000 gram calibration mass.
Tolerance:
The equipment checked shall meet the tolerances specified in the test method listed
above.
Procedure:
1. For calibrating the ignition oven balance, follow the Preventive Maintenance and
Servicing Manual procedure.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #38 Worksheet Ignition
IGNITION OVEN
CERTIFICATE OF VERIFICATION
Procedure No. 38
AASHTO T 308
DATE -
EQUIPMENT VERIFIED -
EQUIPMENT SERIAL NUMBER -
METHOD USED FOR ACCEPTANCE - |  |  | Refer to Operator's Manual
(To Span Calibrate the Balance)
CALIBRATION EQUIPMENT USED -
VERIFICATION WEIGHTS - |  |  |  | 1000.0 gm
| (Read-out on Ignition Oven) |  |  | 3000.0 gm
5000.0 gm
8000.0 gm
EQUIPMENT CHECKED BY -
LOCATION OF EQUIPMENT -
This certificate is issued for the aforementioned equipment and certifies that the equipment has
been checked and found to comply with the requirements set forth in the method listed above.
In order for this equipment to remain in service it must be recertified on the following date and
must comply with the latest methods listed at that time.
NEXT CALIBRATION DUE -
Remarks

## Procedure #37 Gyratory
Procedure for Verifying Gyratory Compactors
Procedure No. 37
Equipment Checked |  |  | GYRATORY COMPACTOR ( AASHTO T 312)
Purpose:
This procedure provides for calibrating the gyratory ram pressure, internal
angle of gyration, gyration frequency and LVDT.
Inspection Equipment Required:
1. Calibration kit, including equipment to verify and calibrate the pressure and angle.
Tolerance:
The equipment checked shall meet the tolerance specified in the equipment owners
manual.
Procedure:
The procedure for calibrating the gyratory compactor is located in the Manual of
Operation and Maintenance for Troxler Model 4140 and 4141. All other gyratories
are calibrated by an outside source and will require certification documents from
those companies.
Verification Interval:
12 months
Report:
A copy of the certificate of calibration must be retained in the Testing Facilities files
for review.

## #37 Worksheet Gyratory
Gyratory Compactor
Procedure No. 37
AASHTO T 312
GYRATORY COMPACTOR CALIBRATION
LABORATORY NAME
LOCATION OF LABORATORY
TYPE OF GYRATORY | [5 empty] | MODEL OF GYRATORY
SERIAL NUMBER | [5 empty] | CALIBRATION KIT USED
CALIBRATION PERFORMED BY | [5 empty] | CALIBRATION KIT LAST CALIBRATED
DAVE-2 CALIBRATION KIT LAST CALIBRATED
COMPANY I.D. TAG # | [5 empty] | COMPACTOR LAST CALIBRATED
1 | PRESSURE CALIBRATION
Date
System Load Cell |  |  |  | kPa |  | TOLERANCE -- |  | 3 Percent or +/- 18 kPa
Calibration Load Cell |  |  |  | kPa
2 | ANGLE CALIBRATION |  |  |  | MOLDS USED
|  | Date |  |  |  |  | TOP | BOTTOM | Avg. | Tolerance
[5 empty] | Internal Angle |  |  |  | 0 | 1.16 +/- .02
Measured Internal Angle |  |  |  | Deg.
[6 empty] | External Angle (If Measured) |  |  |  | Deg.
Ram Head Measurement |  |  | A Reading | B Reading | C Reading |  |  |  |  | 1.25 +/- .02
[6 empty] | Mold Size |  |  | 150 | mm
3 | HEIGHT CALIBRATION
|  | Date |  |  |  | Daily File Kept |  |  |  | y/n
Actual |  |  |  | mm |  | Measured |  |  |  | mm
Pressure |  |  |  | kPa |  | Mold Size |  |  | 150 | mm
4 | ROTATION CALIBRATION |  |  |  | Performed as manufacturers recommendation.
Date
Target |  |  | 30 | RPM |  | Measured |  |  |  | RPM
5 | MAINTENANCE
Follow manufacturers recommended schedule of maintenance. (Includes vacuuming, degreasing, lubricating,
and inspecting for loose or worn parts)
Kept Clean |  |  |  | y/n |  | Well Greased |  |  |  | y/n
OTHER MAINTENANCE PERFORMED -
REMARKS

## Procedure #36 Molds
Procedure for Verifying Gyratory Molds
Procedure No. 36
Equipment Checked: |  | GYRATORY MOLDS AASHTO T 312 -Annex A
Purpose:
This procedure provides instructions for checking the critical dimensions of the molds
and the end plates used in this test method.
Inspection Equipment Required:
1. Calipers or steel rule capable of measuring to 0.01 millimeters.
2. Surface Comparator.
3. Rockwell Hardness Tester.
4. Three Point Bore Gauge or Coordinate Measurement Machine (CMM).
Procedure:
1. Measure and record the inside diameter of the mold to the nearest 0.001 mm. Rotate
the mold in 90 degree intervals, measure and record the inside diameter to get a minimum of
three measurements. Perfom this procedure at the top, bottom and maximum wear position.
2. Measure and record the length of the mold to the nearest millimeter. Repeat this
procedure at intervals of 90 degrees around the mold.
3. Measure the diameter of the mold end plate at intervals of 90 degrees around the plate.
4. Check the smoothness of the molds and bottom plate surfaces that are in contact with
the specimen. Use a surface comparator. (Follow owners manual)
5. Check the Rockwell Hardness of the molds and bottom plates with a Rockwell Hardness
tester. (Follow owners manual)
Note: Only new molds need the hardness tested.
6. Measure and record the diameter gyratory ram at intervals of 90 degrees to the nearest
millimeter.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #36 Worksheet Molds
Gyratory Molds
Procedure No. 36
AASHTO T-312
GYRATORY MOLD CALIBRATION
DATE ----
LOCATION OF EQUIPMENT --
EQUIPMENT CHECKED BY --
METHOD USED FOR ACCEPTANCE -- T-312 Annex A
CALIBRATION EQUIPMENT USED --
New molds inside diameter 149.90 to 150.00 mm. In-service molds inside diameter shall not exceed 150.2 mm.
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
GYRATORY MOLD # |  |  | 1st Reading | 2nd Reading | 3rd Reading | Average | Pass / Fail
| Wall Thickness - (7.5 mm thickness min.) |  |  |  |  | 0
Top | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Bottom | Inside Diameter - (149.90 to 150.20 mm) |  |  |  |  | 0
Middle | Maximum Wear of Mold @ 115 mm |  |  |  |  | 0
| Height of Mold - (250 mm minimum) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
| End Plate - (149.50 to 149.75 mm) |  |  |  |  | 0
Mold #
SURFACE FINISH OF MOLDS-- (1.60 um or smoother)
ROCKWELL HARDNESS (min. C48, One time only)
NEXT CALIBRATION DATE -

## Procedure #23 Vacuum
Procedure for Verifying Vacuum System Pressure
Procedure No. 23
Equipment Checked: |  | VACUUM SYSTEM (AASHTO T 209)
Purpose:
This method provides instructions for checking the pressure developed by vacuum
pumps for the Rice Test.
Procedure:
1. Connect the vacuum pump to the vacuum chamber.
2. Place a manometer in series between the pump and the vacuum chamber being sure
all connections are tight.
3. Start the pump and allow it to run for at least 5 minutes to allow the system to stabilize.
4. Record the vacuum attained and the serial numbers of pump and manometer being used.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #23 Worksheet Vacuum
VERIFYING VACUUM SYSTEM PRESSURE
Procedure No. 23
AASHTO T 209
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
VACUUM SYSTEMS 1-3
System No. | 1 |  |  | 2 |  |  | 3
Manometer Type
Pump Type
No. of Moisture Traps
Micro Valve Installed |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Bowl Calibrated |  |  | Monthly Y/N |  |  | Monthly Y/N |  |  | Monthly Y/N
Wire Mesh on Lid Opening |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Water Bath |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N
Vibrating Shaker Table |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Scale Used (Type)
Vacuum Pressure Obtained (25.0-30.0mm Hg.)
|  |  | mm Hg |  |  | mm Hg |  |  | mm Hg
VACUUM SYSTEMS 4-6
System No. | 4 |  |  | 5 |  |  | 6
Manometer Type
Pump Type
No. of Moisture Traps
Micro Valve Installed |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Bowl Calibrated |  |  | Monthly Y/N |  |  | Monthly Y/N |  |  | Monthly Y/N
Wire Mesh on Lid Opening |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Water Bath |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N |  |  | Heater & Overflow Y/N
Vibrating Shaker Table |  |  | Yes/No |  |  | Yes/No |  |  | Yes/No
Scale Used (Type)
Vacuum Pressure Obtained (25.0-30.0mm Hg.)
|  |  | mm Hg |  |  | mm Hg |  |  | mm Hg
Remarks:

## Procedure #61 Air Meter
Procedure for Calibrating Pressure Air Meter - Type B
Procedure No. 61
Equipment Checked: |  | Pressure Air Meter - Type B (ASTM C 231-10)
Purpose:
This method provides instructions for calibrataing Pressure Air Meter - Type B.
Procedure:
The procedures performed for calibrating Pressure Air Meter shall be in accordance with the
meter type being employed. Refer to ASTM C 231, A1. CALIBRATION OF APPARATUS.
There are varying versions of Pressure Air Meters on the market. It may be necessary to
follow the manufacturers instructions on calibrating the air meter if ASTM procedures do not
apply to meter being calibrated.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #61 Worksheet Air Meter
Pressure Air Meter - Type B
Procedure No. 61
ASTM C 231 - 10
Date: |  | Location of Equipment
Inspected by:
Calibration Equipment Used:
Date: |  |  |  | Date:
Air Meter No.: |  |  |  | Air Meter No.:
Initial Reading: |  |  |  | Initial Reading:
Checked at 5% |  |  |  | Checked at 5%
Checked at 10% |  |  |  | Checked at 10%
Checked at 15% |  |  |  | Checked at 15%
Condition of Meter |  |  |  | Condition of Meter
Date: |  |  |  | Date:
Air Meter No.: |  |  |  | Air Meter No.:
Initial Reading: |  |  |  | Initial Reading:
Checked at 5% |  |  |  | Checked at 5%
Checked at 10% |  |  |  | Checked at 10%
Checked at 15% |  |  |  | Checked at 15%
Condition of Meter |  |  |  | Condition of Meter
Date: |  |  |  | Date:
Air Meter No.: |  |  |  | Air Meter No.:
Initial Reading: |  |  |  | Initial Reading:
Checked at 5% |  |  |  | Checked at 5%
Checked at 10% |  |  |  | Checked at 10%
Checked at 15% |  |  |  | Checked at 15%
Condition of Meter |  |  |  | Condition of Meter
Date: |  |  |  | Date:
Air Meter No.: |  |  |  | Air Meter No.:
Initial Reading: |  |  |  | Initial Reading:
Checked at 5% |  |  |  | Checked at 5%
Checked at 10% |  |  |  | Checked at 10%
Checked at 15% |  |  |  | Checked at 15%
Condition of Meter |  |  |  | Condition of Meter
Date: |  |  |  | Date:
Air Meter No.: |  |  |  | Air Meter No.:
Initial Reading: |  |  |  | Initial Reading:
Checked at 5% |  |  |  | Checked at 5%
Checked at 10% |  |  |  | Checked at 10%
Checked at 15% |  |  |  | Checked at 15%
Condition of Meter |  |  |  | Condition of Meter
Remarks:

## Procedure #11 Bearing Blocks
Procedure for Verifying Bearing Blocks and Retainers
Procedure No. 11
Equipment Checked: |  | BEARING BLOCK (ASTM C 39 M-11a, RETAINERS (ASTM C 1231 M-10a)
Purpose:
This method provides instruction for checking the planeness of the bearing blocks and
retainers of the compression testing machine as described in ASTM C 39 and ASTM C 1231.
Inspection Equipment Required:
1. Steel rule, caliper. Depth gauge (optional)
2. Feeler gauge.
Tolerance:
The critical dimensions shall meet the requirements set forth in ASTM C 39 and ASTM C 1231.
Procedure:
The units shall be checked and any deviation from the requirement shall be repaired prior
to placing the unit back in service.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #11 Worksheet Bearing Blocks
Compression Machine Bearing Block
Planeness and Retainers
Procedure No. 11
ASTM C 39 M-11a & ASTM C 1231 M-10a
Date:
Location of Equipment:
Equipment Serial No.
Checked by:
Calibration Equipment Used:
Upper Bearing Blocks
Diameter (Cylinder Head) |  | 1. |  | 2. | Is face plane? Y / N
Diameter (Cube Head) |  | 1. |  | 2. | Is face plane? Y / N
If not plane, note the amount of wear.
Lower Bearing Blocks
If lower bearing block is round, note the diameter.
Diameter (Cylinder Head) |  | 1. |  | 2. | Is face plane? Y / N
Diameter (Cube Head) |  | 1. |  | 2. | Is face plane? Y / N
If not plane, note the amount of wear.
Retainers (Set of Two)
Retainer Diameter |  |  | Suface Condition |  | (Good, Fair or Poor)
Retainer Diameter |  |  | Suface Condition |  | (Good, Fair or Poor)
Remarks:

## Procedure No. 62 Moist Room
Procedure for Verifying Moist Room & Water Storage Tanks
Procedure No. 62
Equipment Checked: |  | Moist Room & Water Storage Tanks (ASTM C511-09)
Purpose:
This method provides requirements for moist rooms and water storage tanks where
concrete specimens are stored.
Inspection Equipment Required:
1. A calibrated thermometer either Fahrenheit or Celsius graduated in 1.0° increments
having a range which includes the temperature range to be checked.
2. Or a certified Digital temperature measuring device.
3. A sling psychrometer or other moisture measuring device.
Tolerance:
The equipment must meet the tolerances specified in test method ASTM C 511.
Moist Rooms & water storage tanks shall have a temperature of 23.0 ± 2.0 C.
(69.8 - 77.0 F). Moist rooms shall have a humidity of not less than 95%.
Procedure:
Check the temperature of the rooms and tanks with a thermometer or digital
temperature measuring device. Record the results on the verification worksheet.
If a chart or digital recorder is hooked up to or in the moist room or to the tank record that
result also on the worksheet. Check the humidity of the moist room with a sling
psychrometer or other moisture measuring device. Record the result on the verification
worksheet. If a recorder is hooked up to monitor the humidity in the moist room record that
reading also on the verification worksheet. Visually inspect the tank to see if it is
saturated with calcium hydroxide. (11.36g/1gl)
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #62 Worksheet Moist Room
Moisture Room & Water Storage Tank
Conformance Verification
Procedure No. 62
ASTM C 511 - 09
Date:
Inspected by:
Calibration Equipment Used:
Moist Room
| Date | Temperature |  | Humidity |  | Inspector | Comments
| (Monthly) | Recorder | Thermometer | Recorder | Psychrometer
Water Storage Tank
| Date | Temperature |  | Inspector | Comments
| (Monthly) | Recorder | Thermometer

## Procedure #63 Capping Compound
Procedure for Verifying Compressive Strength for Capping Compound
Procedure No. 63
Material Checked: |  | CAPPING COMPOUND (ASTM C617M-11 & ASTM C109M-11a)
Purpose:
This method provides instructions for checking the strength of capping compound.
Inspection Equipment Required:
1. A calibrated concrete compression machine.
2. A heated pot for heating the capping compound.
3. Cube molds.
4. A verified thermometer for checking the temperature of heated capping compound.
5. Straightedge and feeler gauge.
Tolerance:
1. Cylinder strength of 500 to 7000 psi requires 5000 minimum psi for mortar.
2. Cylinder strength above 7000 requires compressive strength not less than the cylinder.
3. Average thickness of the cap at 500 to 7000 psi shall not exceed 1/4 inch.
4. Average thickness of the cap above 7000 psi shall not exceed 1/8 inch.
Procedure:
1. Plug in mortar pot and check molten sulfur contents to verify temperature range of
265 - 290 °F.
2. Stir contents.
3. Lightly coat three cube molds with oil.
4. Using ladle, fill molds with sulfur to the top of filling hole. Wait 15 minutes and refill
to top.
5. After cubes have become hard, strip from molds without breaking off knobs formed
by the filling hole in the cover plate.
6. Remove sharp edges and fins from the cubes. Also remove the oil.
7. Check planeness of cubes to within 0.002 of an inch.
8. Test for strength not less than 24 hours after molding.
9. Test for compressive strength in accordance to C 109.
Verification interval:
Receipt of a new lot
Report:
A copy must be retained in the Testing Facilities files for review.

## #63 Worksheet Capping Compound
Compressive Strength for Capping Compound
Procedure No. 63
ASTM C 617 M-11 & ASTM C 109 M-11 a
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Sulfur Manufacturer
Shipment Date
Date Cube Made
Date Cube Broke
Maximum permissible range between 2 specimens, 7.6% (average)
Maximum permissible range between 3 specimens, 8.7% (average)
Cube #1 Load
Cube #2 Load
Cube #3 Load
| Average |  | 0
Strength average / 4
Remarks:

## Procedure #26 Unit Weight
Procedure for Verifying Unit Weight Measure
Procedure No. 26
Equipment Checked: |  | UNIT WEIGHT MEASURE (ASTM C 138M-106)
Purpose:
This directive gives guidance for the calibration of unit weight measures used in
Aggregate and Portland Cement Concrete Labs.
Inspection Equipment Required:
1. A thermometer capable of being read to the nearest 1.0 degree F.
2. A piece of glass large enough to cover the entire top of the measure being calibrated.
3. A scale capable of weighing the yield measure, glass, and the water within to the
nearest 0.01 lbs.
Procedure:
Follow the procedure set forth in ASTM C 29, Section 8.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #26 Worksheet Unit Weight
Unit Weight Measure
Procedure No. 26
AASHTO C 138M-106
Date: |  |  |  | Location of Equipment
Checked by: | [5 empty] | Equipment Serial No.:
CALIBRATING UNIT WEIGHT MEASURE
TEMPERATURE CHART
Wt of Bowl-Glass Plate with Water | [6 empty] | C | F | kg/m3 | lb/ft3
[7 empty] | 15.6 | 60 | 999.01 | 62.366
Wt of Bowl-Glass Plate | [6 empty] | 18.3 | 65 | 998.54 | 62.336
[7 empty] | 21.1 | 70 | 997.97 | 62.301
Net Mass of Water, (nearest lb.) |  |  |  | (M) |  | (M) | 23 | 73.4 | 997.54 | 62.274
[7 empty] | 23.9 | 75 | 997.32 | 62.261
Density of H2O kg/m3 (Take off Chart) |  |  |  | (D) |  | (D) | 26.7 | 80 | 996.59 | 62.216
[7 empty] | 29.4 | 85 | 995.83 | 62.166
Volume of Unit Weight nearest (lb.) |  |  |  | lb. (V) |  | lb. (V)
[6 empty] | 1 | 2
Calculation | V=1000 | M |  |  | Avg. of 2
|  | D |  |  |  |  | Avg of 2
Weigh the measure and glass plate. Fill the measure with deionized water at a
temperature of 18 to 24 degree C. Record the temperature of the water. Slide the glass plate on the
measure, being sure no air bubbles remain. Dry the outer surface of the measure and determine
the combined mass of measure, glass plate, and water by weighing. Following the final
weighing determine the mass of the clean,dry,empty measure for subsequent tests.
AASHTO T 304-5 (8)
Bowl Number
Weight of Bowl
Volume of Bowl
BOWL DIMENSIONS
Outside diameter
Inside diameter
Inside height
Outside height

## Procedure #65 Vol. Air Meter
Procedure for Calibrating Volumetric Air Meter
Procedure No. 65
Equipment Checked: |  | Volumetric Air Meter (ASTM C 173-08)
Purpose:
This method provides instructions for calibrating a Volumetric Air Meter.
Inspection Equipment Required:
1. Thermometer
2. Glass Plate
3. Scale
Procedure:
1. Determine the volume of the bowl with an accuracy of at least 0.1 % by determining the
mass of water required to fill it at room temperature and dividing it by the density of
the water at the same temperature. See ASTM C 29
2. Determine the accuracy of the graduations on the neck of the top section of the air meter
by filling the assembled measuring bowl and top section with water to the level of the mark
for highest air content graduation.
3. Add water in increments of 1.0% of the volume of the bowl to check the accuracy
throughout the graduated range of air content. The error at any point throughout the
graduated range shall not exceed 0.1% of air.
4. Determine the volume of the calibrated cup using water at 70°F by the method outlined in
Step 1.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #65 Worksheet Vol. Air Meter
Volumetric Air Meter
Procedure No.65
ASTM C 173-08
Date:
Inspected by:
Air Meter No.:
Calibration Equipment Required:
Volume of Bowl
Mass of the water, glass plate & bowl: |  |  |  |  | (W)
Mass of glass plate & bowl: |  |  |  |  | (M)
Density of Water @ 70°F (62.3 lb/ft3): |  |  |  |  | (D)
Volume: |  | V=(W-M)/D
(Bowl capacity is greater than 0.075 cubic feet (2L))
Volume of Cup
Mass of the water, glass plate & cup: |  |  |  |  | (W)
Mass of glass plate & cup: |  |  |  |  | (M)
Density of Water @ 70°F (62.3 lb/ft3): |  |  |  |  | (D)
Volume: |  | V=(W-M)/D
(Capacity of or being graduated in increments equal to 1.00 ± .04% of the bowl volume)
Calibration of Volumetric Air Meter Neck
Actual Reading @
9% | 8% | 7% | 6% | 5% | 4% | 3% | 2% | 1% | 0%
(Graduations accurate to ± 0.1% by volume of the bowl)
Comments:

## Procedure #18 Liquid Limit
Procedure for Verifying Liquid Limit Devices And Grooving Tools
Procedure No. 18
Equipment Checked: |  | LIQUID LIMIT DEVICE AND GROOVING TOOL (AASHTO T 89)
Purpose:
This method provides instructions for checking the liquid limit device and grooving
tool used in this test.
Inspection Equipment Required:
1. Caliper capable of reading to nearest 0.01 mm.
2. Dial micrometer.
Tolerance:
The critical dimensions shall meet the requirements set forth in AASHTO T 89.
Procedure:
1. Check all the critical dimensions for the liquid limit device's brass cup and the grooving
tool. A picture of the apparatus and the necessary points to measure with the
tolerances are printed in the AASHTO test procedure T 89.
2. Where the cup comes in contact with the base, a small indention will be made.
Measure the depth of this indention using the vertical dial micrometer to the nearest
0.01 in.
3. Check the rim of the cup for excessive wear, which will cause it to become sharp
rather than a milled flat surface. Rejection is at the inspector's judgment.
4. Turn the crank slowly and note if the cup wobbles on the rise. Also check that the
cup is fastened to the cam follower with mechanical rather than a welded connector.
5. Check the inside of the cup for excessive wear, which will create a groove in the cup.
Rejection is based on the judgment of the inspector.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities file for review.

## #18 Worksheet Liquid Limit
Liquid Limit Equipment
Procedure No. 18
AASHTO T 89
Date:
Checked by:
Location of Equipment:
Calibration Equipment Used:
Liquid Limit Device
|  | Cup Assembly |  |  |  | Base
Dimension |  | A | B | C | N | K | L | M
Description |  | Radius of Cup | Thickness of Cup | Depth of Cup | Cup at Cam Follower to Base | Thickness | Length | Width
Metric, mm |  | 54 | 2 | 27 | 47 | 50 | 150 | 125
Tolerance, mm |  | 2 | 0.1 | 1 | 1.5 | 5 | 5 | 5
Lab Equip. Number
Lab Equip. Number
Lab Equip. Number
Does cup wobble on rise? |  |  | 1. |  | Cup groove condition? |  |  | 1.
(Yes or No) |  |  | 2. |  | (Good or Bad) |  |  | 2.
|  |  | 3. |  |  |  |  | 3.
Note: | Plate "H" may be designed for using one securing screw.
Remarks:

## Grooving Tool T 89
Liquid Limit Equipment
AASHTO T 89
Grooving Tool
|  |  |  | Curved End |  |  | Gage
|  | Dimensions |  | a | b | c | d | e
|  | Description |  | Thickness | Cutting Edge | Width | Depth | Length
|  | Metric, mm |  | 10 | 2 | 13.5 | 10 | 15.9
|  | Tolerance, mm |  | 0.1 | 0.1 | 0.1 | 0.2 | NA
Lab Equip. Number
Lab Equip. Number
Lab Equip. Number
Note: | An additional wear tolerance of 0.1 mm shall be allowed for dimension "b" for used grooving tools.
Remarks:

## Grooving Tool D 4318
Liquid Limit Equipment
ASTM D 4318
Grooving Tool
| Letter | AΔ | BΔ | CΔ | DΔ | EΔ | FΔ
| MM | 2 + 0.1 | 11 + 0.2 | 40 + 0.5 | 8 + 0.1 | 50 + 0.5 | 2 + 0.1
Lab Equip. No.
Lab Equip. No.
Lab Equip. No.
| Letter | G | H | J | KΔ | LΔ | N
| MM | 10 Min. | 13 | 60 | 10 + 0.05 | 60 Deg + 1 | 20
Lab Equip. No.
Lab Equip. No.
Lab Equip. No.
Remarks:

## Procedure #5 & 5a Manual Hammer
Procedure for Verifying Manual Compaction Hammers
Procedure No. 5 & 5a
Equipment Checked: |  | MANUAL HAMMER (AASHTO T 99 & T 180)
Purpose:
This method provides instructions for checking the critical dimensions of the proctor
hammer.
Inspection Equipment Used:
1. Calipers readable to 0.001 inch.
2. Tape measure readable to 1/16 inch.
3. Balance, capacity 5 kg., readable to 1 gram.
Tolerance:
Equipment shall meet the dimensional tolerances specified in the above test.
Procedure:
1. Measure and record the diameter of the rammer face determined by taking two readings
90 degrees apart using the calipers.
2. Pull up the handle, measure and record the drop height of the hammer. Determine this
height inside the guide sleeve using the tape measure.
3. Remove the hammer from the guide sleeve. Determine and record its mass to the
nearest 1 gram.
4. Measure and record the diameters of the vent holes near the end of the hammer.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #5 Worksheet Manual Rammer
Manual Rammer
Procedure No. 5
AASHTO T 99
Date:
Location of Equipment:
Hammer Number:
Checked By:
Calibration Equipment Used:
Manual Rammer
Is the face flat? |  |  | Yes | No
Diameter of the face |  |  | Dia. # 1 | Dia. # 2 | ( > 1.985 in.)
Drop height of the rammer |  |  |  | (12 + 0.06 in.)
Rammer Mass |  |  |  | (2495 + 9 grams)
Diameter of Vent Holes
Top # 1 |  |  | Top # 2 | Top # 3 | Top # 4
Bott. # 1 |  |  | Bott. # 2 | Bott. # 3 | Bott # 4
Specification: |  | Min. 9.5mm or 3/8 in. Dia.
Note: Height may be impractical to adjust. Adjustments would have to be made with
the height of the material that is being compacted.
Remarks:

## #5a Worksheet Modified Manual
Manual Rammer
Procedure No. 5a
AASHTO T 180
Date:
Location of Equipment:
Hammer Number:
Checked by:
Calibration Equipment:
Manual Rammer
Is the face of the rammer flat? |  |  | Yes |  | No
Diameter of the face. |  |  | Dia. # 1 |  | Dia. # 2 |  | (> 1.985 in.)
Drop height of rammer. |  |  |  | (18 + 0.06 in.)
Rammer mass. |  |  |  | (4536 + 9 grams)
Diameter of Vent Holes
Top # 1 |  | Top # 2 |  | Top # 3 |  | Top # 4
Bott. # 1 |  | Bott. # 2 |  | Bott. # 3 |  | Bott. # 4
Specification: |  | Min. 9.5mm or 3/8 in. Dia.
Note: Height may be impractical to adjust. Adjustments would have to be made with
the height of the material that is being compacted.
Remarks:

## Procedure #22 & 22a Mechanical
Procedure for Verifying Mechanical Soil Rammers
Procedure No. 22 & 22a
Equipment Checked: |  | MECHANICAL SOIL RAMMER (AASHTO T 99)(AASHTO T 180)
Purpose:
This method provides instructions for checking the critical dimensions and masses
required in the test method for this piece of equipment
Inspection Equipment Required:
1. Outside caliper with a 2 in. capacity readable to 0.001 in.
2. A hand rammer meeting the requirements of the AASHTO Test Method being
checked (T 99 or T 180)
3. A scale meeting AASHTO M 231 for Class G-2 and Class G-20.
4. A drying oven capable of maintaining a temperature not to exceed 140o F.
Tolerance:
The dimensions and masses shall meet the requirements set forth in the AASHTO Test
Method being checked (T 99 or T 180).
Procedure:
1. Use the caliper to measure the diameter of the rammer face at two locations 90
degrees apart and record.
2. Using the edge of the ruler check to see if the face of the rammer I s flat and record.
3. Calibrate the mechanical rammer by using the procedure set forth I n ASTM D 2168.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #22 Worksheet T99 Mechanical
Mechanical Soil Rammer
Procedure No. 22
AASHTO T 99
Date:
Location of Equipment:
Hammer Number:
Checked by:
Calibration Equipment Used:
Mechanical Rammer (AASHTO T 99)
Is the face of the rammer flat? |  |  |  | Yes | No
Diameter of the face. |  |  |  | Dia. # 1 | Dia. # 2 | (> 1.985 in.)
Drop height of rammer. |  |  |  |  | (12 + 0.06 in.)
Rammer mass. |  |  |  |  | (2495 + 9 grams)
Remarks:

## #22 Worksheet T99-1 Mechanical
Calibration of Laboratory Mechanical Rammers
Procedure No. 22-1 (for Standard Proctor ASSHTO T 99
ASTM D2168
Date:
Equipment Make:
Model Number:
Serial Number:
Manual Rammer Equipment Number Used:
Calibration Practice: A or B
Practice A Calculations:
= ( gmax ) the maximum dry unit weight obtained with the manual compactor
= ( g'max ) the maximum dry unit weight obtained with the mechanical compactor
= ( W ) the percentage difference of maximum dry unit weight values
W = [(g'max - gmax) / gmax ] X 100
* | If the absolute value of W is equal to or less than 2.0, the mechanical compactor
is satisfactory for immediate use. If the absolute value of W is greater than 2.0,
then follow the procedures detailed in ASTM D2168.
Practice B Calculations:
* | Follow procedure for Practice B detailed in ASTM D2168
= ( c1 ) micrometer reading before impact, manual rammer
= ( c2 ) micrometer reading after impact, manual rammer
= ( c'1 ) micrometer reading before impact, mechanical rammer
= ( c'2 ) micrometer reading after impact, mechanical rammer
= ( D ) deformation value for single lead cylinder, manual rammer
= ( D' ) deformation value for single lead cylinder, mechanical rammer
= ( Dave and D'ave) average deformation values where v1 & v'1 values do not exceed +2.0%
= ( v1 ) % variation from the mean of individual deformation values, manual rammer
= ( v'1 ) % variation from the mean of individual deformation values, mechanical rammer
= ( v2 ) % difference of the average deformation value of the mechanical vs. manual rammer
= ( N ) number of tests
| D = c1 - c2 |  |  | D' = c'1 - c'2
| Dave = S ( c1 - c2 ) / N |  |  | D'ave = S ( c'1 - c'2 ) / N
| v1 = [ (D - Dave) / Dave ] X 100 |  |  | v'1 = [ (D' - D'ave) / D'ave ] X 100
v2 = [ (Dave - D'ave) / Dave] X 100
Remarks / Adjustments:

## #22a Worksheet T180 Mod. Mech.
Mechanical Soil Rammer
Procedure No. 22a
AASHTO T 180
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Mechanical Rammer #1 (AASHTO T 180)
Hammer Number:
Is the face of the rammer flat? |  |  |  | Yes | No
Diameter of the face. |  |  |  | Dia. # 1 | Dia. # 2 | (> 1.985 in.)
Drop height of rammer. |  |  |  |  | (18 + 0.06 in.)
Rammer mass. |  |  |  |  | (4536 + 9 grams)
Mechanical Rammer #2 (AASHTO T 180)
Hammer Number:
Is the face of the rammer flat? |  |  |  | Yes | No
Diameter of the face. |  |  |  | Dia. # 1 | Dia. # 2 | (> 1.985 in.)
Drop height of rammer. |  |  |  |  | (18 + 0.06 in.)
Rammer mass. |  |  |  |  | (4536 + 9 grams)
Mechanical Rammer #3 (AASHTO T 180)
Hammer Number:
Is the face of the rammer flat? |  |  |  | Yes | No
Diameter of the face. |  |  |  | Dia. # 1 | Dia. # 2 | (> 1.985 in.)
Drop height of rammer. |  |  |  |  | (18 + 0.06 in.)
Rammer mass. |  |  |  |  | (4536 + 9 grams)
Remarks:

## #22a Worksheet T180-1 Mod. Mech
Calibration of Laboratory Mechanical Rammers
Procedure No. 22a-1 (for Modified Proctor ASSHTO T 180
ASTM D2168
Date:
Equipment Make:
Model Number:
Serial Number:
Manual Rammer Equipment Number Used:
Calibration Practice: A or B
Practice A Calculations:
= ( gmax ) the maximum dry unit weight obtained with the manual compactor
= ( g'max ) the maximum dry unit weight obtained with the mechanical compactor
= ( W ) the percentage difference of maximum dry unit weight values
W = [(g'max - gmax) / gmax ] X 100
* | If the absolute value of W is equal to or less than 2.0, the mechanical compactor
is satisfactory for immediate use. If the absolute value of W is greater than 2.0,
then follow the procedures detailed in ASTM D2168.
Practice B Calculations:
* | Follow procedure for Practice B detailed in ASTM D2168
= ( c1 ) micrometer reading before impact, manual rammer
= ( c2 ) micrometer reading after impact, manual rammer
= ( c'1 ) micrometer reading before impact, mechanical rammer
= ( c'2 ) micrometer reading after impact, mechanical rammer
= ( D ) deformation value for single lead cylinder, manual rammer
= ( D' ) deformation value for single lead cylinder, mechanical rammer
= ( Dave and D'ave) average deformation values where v1 & v'1 values do not exceed +2.0%
= ( v1 ) % variation from the mean of individual deformation values, manual rammer
= ( v'1 ) % variation from the mean of individual deformation values, mechanical rammer
= ( v2 ) % difference of the average deformation value of the mechanical vs. manual rammer
= ( N ) number of tests
| D = c1 - c2 |  |  | D' = c'1 - c'2
| Dave = S ( c1 - c2 ) / N |  |  | D'ave = S ( c'1 - c'2 ) / N
| v1 = [ (D - Dave) / Dave ] X 100 |  |  | v'1 = [ (D' - D'ave) / D'ave ] X 100
v2 = [ (Dave - D'ave) / Dave] X 100
Remarks / Adjustments:

## Procedure #2 Comp Molds
Date: | 2-3-97
Revised:
Procedure for Verifying Compaction Molds
Procedure No. 2
Equipment Checked: |  | COMPACTION MOLDS (AASHTO T 99, T 180)
Purpose:
This method provides instructions for checking the critical dimensions of the molds
used in this test.
Inspection Equipment Required:
1. Calipers capable of measuring an inside diameter of 4 in. and readable to 0.001 in.
2. Calipers capable of measuring an outside height of 6 in. and readable to 0.001 in.
checked (T 99 or T 180)
Tolerance:
Molds shall meet all dimension requirements set forth in the applicable test method
listed above.
Procedure:
1. Measure and record the inside diameter of the mold to nearest 0.001 in. Rotate
mold 90 degrees and repeat the procedure. Turn the mold over and repeat.
2. Measure and record the height of the mold less its collars to the nearest 0.001 in.
Repeat this procedure at intervals of 90 degrees around the mold and record.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #2 Worksheet-4" Comp Molds
Compaction Mold
Procedure No. 2
AASHTO T 99 & T 180
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Mold Measurements
4 inch Molds |  |  |  | 4 inch Molds
Diameter, Inches |  | Height, Inches |  | Diameter, Inches | Height, Inches
( 4.000 + 0.016 ) |  | ( 4.584 + 0.018 ) |  | ( 4.000 + 0.016 ) | ( 4.584 + 0.018)
Dia. # 1 |  | Ht. # 1 |  | Dia. # 1 | Ht. # 1
Dia. # 2 |  | Ht. # 2 |  | Dia. # 2 | Ht. # 2
Dia. # 3 |  | Ht. # 3 |  | Dia. # 3 | Ht. # 3
Dia. # 4 |  | Ht. # 4 |  | Dia. # 4 | Ht. # 4
Ave. Diameter, in. |  | Ave. Height, in. |  | Ave. Diameter, in. | Ave. Height, in.
0 |  | 0 |  | 0 | 0
Mold Number: |  |  |  | Mold Number:
Computed Volume, cu, Inches |  |  | 0 | Computed Volume, cu, Inches |  | 0
Computed Volume, cu, Feet |  |  | 0 | Computed Volume, cu, Feet |  | 0
Computed Volume, cu Cm. |  |  | 0 | Computed Volume, cu Cm. |  | 0
Mold Measurements
4 inch Molds |  |  |  | 4 inch Molds
Diameter, Inches |  | Height, Inches |  | Diameter, Inches | Height, Inches
( 4.000 + 0.016 ) |  | ( 4.584 + 0.018) |  | ( 4.000 + 0.016 ) | ( 4.584 + 0.018)
Dia. # 1 |  | Ht. # 1 |  | Dia. # 1 | Ht. # 1
Dia. # 2 |  | Ht. # 2 |  | Dia. # 2 | Ht. # 2
Dia. # 3 |  | Ht. # 3 |  | Dia. # 3 | Ht. # 3
Dia. # 4 |  | Ht. # 4 |  | Dia. # 4 | Ht. # 4
Ave. Diameter, in. |  | Ave. Height, in. |  | Ave. Diameter, in. | Ave. Height, in.
0 |  | 0 |  | 0 | 0
Mold Number: |  |  |  | Mold Number:
Computed Volume, cu, Inches |  |  | 0 | Computed Volume, cu, Inches |  | 0
Computed Volume, cu, Feet |  |  | 0 | Computed Volume, cu, Feet |  | 0
Computed Volume, cu Cm. |  |  | 0 | Computed Volume, cu Cm. |  | 0
Note: | A mold that fails to meet manufacturing tolerances after continued service may remain
in use provided those tolerances are not exceeded by more that 50 percent; and the volume
of the mold, calibrated in accordance with T 19M/T 19 for Unit Mass of Aggregate, is used in
the calculations.
Remarks:

## #2 Worksheet-6" Comp Molds
Compaction Mold
Procedure No. 2
AASHTO T 99 & T 180
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Mold Measurements
6 inch Molds |  |  |  | 6 inch Molds
Diameter, Inches |  | Height, inches |  | Diameter, Inches | Height, inches
( 6.000 + 0.026 ) |  | ( 4.584 + 0.018) |  | ( 6.000 + 0.026 ) | ( 4.584 + 0.018)
Dia. # 1 |  | Ht. # 1 |  | Dia. # 1 | Ht. # 1
Dia. # 2 |  | Ht. # 2 |  | Dia. # 2 | Ht. # 2
Dia. # 3 |  | Ht. # 3 |  | Dia. # 3 | Ht. # 3
Dia. # 4 |  | Ht. # 4 |  | Dia. # 4 | Ht. # 4
Ave. Diameter, in. |  | Ave. Height, in. |  | Ave. Diameter, in. | Ave. Height, in.
0 |  | 0 |  | 0 | 0
Mold Number: |  |  |  | Mold Number:
Computed Volume, cu, Inches |  |  | 0 | Computed Volume, cu, Inches |  | 0
Computed Volume, cu, Feet |  |  | 0 | Computed Volume, cu, Feet |  | 0
Computed Volume, cu Cm. |  |  | 0 | Computed Volume, cu Cm. |  | 0
Mold Measurements
6 inch Molds |  |  |  | 6 inch Molds
Diameter, Inches |  | Height, inches |  | Diameter, Inches | Height, inches
( 6.000 + 0.026 ) |  | ( 4.584 + 0.018) |  | ( 6.000 + 0.026 ) | ( 4.584 + 0.018)
Dia. # 1 |  | Ht. # 1 |  | Dia. # 1 | Ht. # 1
Dia. # 2 |  | Ht. # 2 |  | Dia. # 2 | Ht. # 2
Dia. # 3 |  | Ht. # 3 |  | Dia. # 3 | Ht. # 3
Dia. # 4 |  | Ht. # 4 |  | Dia. # 4 | Ht. # 4
Ave. Diameter, in. |  | Ave. Height, in. |  | Ave. Diameter, in. | Ave. Height, in.
0 |  | 0 |  | 0 | 0
Mold Number: |  |  |  | Mold Number:
Computed Volume, cu, Inches |  |  | 0 | Computed Volume, cu, Inches |  | 0
Computed Volume, cu, Feet |  |  | 0 | Computed Volume, cu, Feet |  | 0
Computed Volume, cu Cm. |  |  | 0 | Computed Volume, cu Cm. |  | 0
Note: | A mold that fails to meet manufacturing tolerances after continued service may remain
in use provided those tolerances are not exceeded by more that 50 percent; and the volume
of the mold, calibrated in accordance with T 19M/T 19 for Unit Mass of Aggregate, is used in
the calculations.
Remarks:

## Procedure #19 straightedge
Procedure for Verifying Straightedges
Procedure No. 19
Equipment Checked: |  | STRAIGHTEDGE (AASHTO T 99, T 180)
Purpose:
This method provides instructions for checking the steel straightedge used in AASHTO
Test Methods T 99 and T 180.
Inspection Equipment Required:
1. Ruler.
2. Feeler gauge.
Tolerance:
The critical dimensions shall meet the requirements set forth in AASHTO T 99 and T 180.
Procedure:
1. Measure the length with a ruler and record.
2. Place the straightedge so that the beveled edge is down on a flat surface and using
a 0.01 in. thickness gauge, try to pass the gauge between the edge of the unit and the
flat surface. If the gauge can be slipped between the unit and the flat surface the
straightedge shall have to be repaired.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #19 Worksheet Straightedge
STEEL STRAIGHTEDGE
Procedure No. 19
AASHTO T 99 , T 180
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Steel Straightedge - #1
Serial Number:
Is unit made of hardened steel? |  |  | Yes | No
Is the unit at least 10 inches in length? |  |  | Yes | No
Is one edge beveled? |  |  | Yes | No
Is the edge plane to within 0.01 inch per 10 inches of length? |  |  |  | Yes | No
Steel Straightedge - #2
Serial Number:
Is unit made of hardened steel? |  |  | Yes | No
Is the unit at least 10 inches in length? |  |  | Yes | No
Is one edge beveled? |  |  | Yes | No
Is the edge plane to within 0.01 inch per 10 inches of length? |  |  |  | Yes | No
Steel Straightedge - #3
Serial Number:
Is unit made of hardened steel? |  |  | Yes | No
Is the unit at least 10 inches in length? |  |  | Yes | No
Is one edge beveled? |  |  | Yes | No
Is the edge plane to within 0.01 inch per 10 inches of length? |  |  |  | Yes | No
Steel Straightedge - #4
Serial Number:
Is unit made of hardened steel? |  |  | Yes | No
Is the unit at least 10 inches in length? |  |  | Yes | No
Is one edge beveled? |  |  | Yes | No
Is the edge plane to within 0.01 inch per 10 inches of length? |  |  |  | Yes | No
Remarks:

## Procedure #8 Balance & Weights
Procedure for Verifying Scales, Balances and Balance Weights
Procedure No. 8
Equipment Checked: |  | SCALES AND BALANCES (AASHTO M231)
Purpose:
This method provides instructions for checking the accuracy sensitivity of scales, balances
balance weights used in all labs.
Inspection Equipment Required:
1. Calibrated weights with certificates.
2. Powder free gloves designed for handling these weights.
Tolerance:
All AASHTO and ASTM tests requiring the use of scales and or balances have in their
procedures a list of requirements for the weighing devices. It will be up to the inspector
to check these requirements when checking and verifying this equipment.
Procedure:
Each balance or scale will have its own worksheet. Choose the appropriate worksheet for
the balance or scale to be verified. Check the equipment through the entire range on the
worksheet. It is recommended to calibrate the balance before verifying. Follow the owners
manual for this procedure.
Lab balance weights used only for verification purposes shall be verified against calibrated
weights.
Verification Interval:
12 months
Report:
A copy for each scale or balance must be retained in the Testing Facilities files for review.

## #8 Balance Weights Worksheet
Balance Weights
Procedure No. 8
AASHTO M 231 Table 4
Date :
Serial No.:
Checked by:
Material:
Scale No.:
Calibration Equipment:
Balance Weights
Weight | Accuracy
5000g | .80g
2000g | .40g
1000g | .25g
500g | .20g
200g | .10g
100g | .07g
50g | .04g
20g | .02g
10g | .015g
5g | .010g
2g | .006g
1g | .004g
Remarks:

## #8 Analytical Worksheet
Analytical Balances
Procedure No. 8
1200g Capacity or Less
Class C
Date:
Inspected By:
Calibration Weights:
Scale Serial No.:
Manufacturer:
Scale Capacity:
Test Point |  | Accuracy Required |  | Accuracy
50 g |  | .02 g
100 g |  | .02 g
200 g |  | .02 g
300 g |  | .02 g
400 g |  | .02 g
500 g |  | .02 g
600 g |  | .02 g
700 g |  | .02 g
800 g |  | .02 g
900 g |  | .02 g
1000 g |  | .02 g
1100 g |  | .02 g
1200 g |  | .02 g
Remarks:

## #8 Gen. Purp. Worksheet
General Purpose Scales
Procedure No. 8
Class G2
Date:
Location of Scale:
Inspected By:
Calibration Weights:
Scale No.:
Manufacturer:
Scale Capacity:
Test Point |  | Accuracy Required | Accuracy
50 g |  | .2 g
100 g |  | .2 g
200 g |  | .2 g
300 g |  | .3 g
400 g |  | .4 g
500 g |  | .5 g
1000 g |  | 1.0 g
2000 g |  | 2.0 g
3000 g |  | 3.0 g
4000 g |  | 4.0 g
5000 g |  | 5.0 g
6000 g |  | 6.0 g
7000 g |  | 7.0 g
8000 g |  | 8.0 g
9000 g |  | 9.0 g
10000 g |  | 10.0 g
11000 g |  | 11.0 g
12000 g |  | 12.0 g
13000 g |  | 13.0 g
14000 g |  | 14.0 g
15000 g |  | 15.0 g
16000 g |  | 16.0 g
Remarks:

## #8 Gen. Purp. > 40lb Worksheet
General Purpose Scales
Procedure No. 8
Class G300 Special (Over 40 LB)
Date:
Location of Scale:
Inspected By:
SN of Calibration Weights:
Scale No.:
Manufacturer:
Scale Capacity:
Test Point |  | Accuracy Required | Accuracy
10 lb |  | .5 lb
20 lb |  | .5 lb
30 lb |  | .5 lb
40 lb |  | .5 lb
50 lb |  | .5 lb
60 lb |  | .6 lb
90 lb |  | .9 lb
120 lb |  | 1.2 lb
150 lb |  | 1.5 lb
180 lb |  | 1.8 lb
210 lb |  | 2.1 lb
240 lb |  | 2.4 lb
270 lb |  | 2.7 lb
300 lb |  | 3.0 lb
Remarks:

## Procedure #1 Oven
Procedure for Verifying Ovens
Procedure No. 1
Equipment Checked: |  | DRYING OVENS
Purpose:
This method provides instructions for checking drying ovens used in the laboratory.
Inspection Equipment Required:
1. A calibrated thermometer either Fahrenheit or Celsius graduated in 1.0° increments
having a range which includes the temperature range to be checked.
2. A thermometer well to retain heat while the oven door is open.
3. A clothes pin the hold the thermometer in such a manner as to enable the operator
to read the scale easily.
Tolerance:
Drying ovens shall be capable of maintaining a constant temperature range listed in the
appropriate test methods.
Procedure:
1. Place the thermometer inside the well with the clothes pin attached to the
thermometer. Position the thermometer on the shelf where the samples are normally
dried.
2. Take the first reading at least 1 hour after closing the oven (oven should remain
undisturbed).
3. Take as many readings as necessary to determine if the temperature range is within
the specified tolerance (three consecutive readings, taken no less that 1/2 hour apart,
within the tolerance allowed are adequate).
4. Adjust the temperature of the oven if an observed reading is outside the tolerance
specified (allow at least 1/2 hr. for the temperature to stabilize between each
adjustment. Return to step 3.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #1 Worksheets Oven
Ovens
Procedure No. 1
AASHTO T-208, T-216, T-265, T-296
Date - |  |  | Inspected by:
Location of Equipment -
Calibration Equipment Used -
Oven #1 |  |  |  | Oven #2
Manufacturer |  |  |  | Manufacturer
Model No. |  |  |  | Model No.
Serial No. |  |  |  | Serial No.
Oven Set Point |  | Temperature Reading |  | Oven Set Point |  | Temperature Reading
| F | 20 min. |  |  | F | 20 min.
|  | 40 min. |  |  |  | 40 min.
|  | 60 min. |  |  |  | 60 min.
( 230+/- 9 Degrees F or 110 +/- 5 Degrees C) |  |  |  | ( 230+/- 9 Degrees F or 110 +/- 5 Degrees C)
Remarks |  |  |  | Remarks
Oven #3 |  |  |  | Oven #4
Manufacturer |  |  |  | Manufacturer
Model No. |  |  |  | Model No.
Serial No. |  |  |  | Serial No.
Oven Set Point |  | Temperature Reading |  | Oven Set Point |  | Temperature Reading
| F | 20 min. |  |  | F | 20 min.
|  | 40 min. |  |  |  | 40 min.
|  | 60 min. |  |  |  | 60 min.
( 230+/- 9 Degrees F or 110 +/- 5 Degrees C) |  |  |  | ( 230+/- 9 Degrees F or 110 +/- 5 Degrees C)
Remarks |  |  |  | Remarks
Remarks

## Procedure #27 Shaker
Procedure for Verifying Mechanical Sieve Shakers
Procedure No. 27
Equipment Checked: |  | MECHCANICAL SIEVE SHAKER (AASHTO T 27)
Purpose:
This method provides instructions for checking the sieving efficiency of mechanical
sieve shakers.
Inspection Equipment Required:
1. A scale capable of weighing to the nearest 0.1 gram.
2. An oven for drying the sample to a constant weight.
3. Sieves of sufficient diameter and screen openings so as not to cause overloading of
any one screen.
4. Sample splitter.
Procedure:
1. Dry the sample to a constant weight.
2. Split the dried sample so as to provide at least 500 grams of fine aggregate or 2500 grams
of coarse aggregate.
3. Place the prepared and weighed sample in the sieve nest and place on the shaker.
4. Shake the sample for 5 minutes and remove.
5. Hand shake each sieve over a pan for one minute and weigh the material passing and
record the weight for each sieve,
6. The material passing from hand shaking in Step 5 shall not exceed 0.5% of the total
weight of the sample used.
7. If the weight exceeds the requirement in #6, increase the time on the shaker and repeat
the test until the requirement is met.
8. When checking shaker which have adjustable shaking speeds it may be necessary to
adjust the speed as well as the time in order to meet the requirements.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #27 Worksheet Shaker
Mechanical Sieve Shakers
Sieving Thoroughness
Procedure No. 27
AASHTO T 27
Date:
Checked by:
Shaker Serial No.:
Manufacturer:
Type of Aggregate used: |  |  | Fine |  | Coarse
Scale Serial Number:
Oven Serial Number:
Sieve Diameter: |  | 8" |  | 10" |  | 12"
Initial Sample Weight: |  |  |  | gm.
Sieve |  | Screen |  | Material |  |  | Percent
Serial No. |  | No. |  | Passing |  |  | Passing
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
[5 empty] | gm. |  |  | %
Total Mechanical Shaker Time |  |  |  | min.
NOTE: Use a new worksheet for each run and or type of aggregate used.

## Procedure #25 Sieve
Procedure for Verifying Sieves
Procedure No. 25
Equipment Checked: |  | SIEVES (ASTM E-11)
Purpose:
This method provides instructions for visually checking sieves used in the laboratory.
Inspection Equipment Required:
1. A method for providing a lighted background when viewing the sieve screen.
2. A magnifier for close examination of the sieve screen.
Tolerance:
The sieves examined shall meet the requirements set forth in ASTM E-11.
Procedure:
1. Check each sieve tag for the proper designation and standard opening size. This
designation shall be in metric dimensions and in the case of older sieves the English
and metric dimensions shall be both on the tag.
2. Check the frame for excessive wear along the rim and for looseness between the top
and bottom halves.
3. Check the screen for loose wires, bowed screens, torn screens, creases I n the screen,
broken screen wires, and deformation of the screen openings. These items are normally
visible with the naked eye when viewed against a lighted background. The use of a
magnifier will aid when viewing the finer screens.
Verification Interval:
12 months
Report:
A copy must be retained in the Testing Facilities files for review.

## #25 Worksheet Sieve
Wire Cloth Sieves for Testing Purposes
Procedure No. 25
ASTM E-11
Date:
Location of Equipment:
Checked by:
Calibration Equipment Used:
Sieves
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Lab Identification
Diameter of Frame
Screen No.
Frame Condition
Screen Condition
Label Condition
Remarks:

## Procedure #64 Thermometer
Procedure for Verifying Thermometers
Procedure No. 64
Equipment Checked: |  | Thermometers
Purpose:
This method provides instructions for checking thermometers used in the laboratory and in
the field.
Inspection Equipment Required:
1. A calibrated thermometer either Fahrenheit or Celsius graduated in 1.0° increments
having a range which includes the temperature range to be checked. A calibrated
digital temperature measuring device is also acceptable.
Tolerance:
The tolerance for mercury thermometers is plus or minus one scale division of that
thermometer. Tolerance for spirit filled thermometers is plus or minus 1% of the total range
for that thermometer. Tolerance for dial thermometers is plus or minus one scale division.
Digital thermometers such as Oaktrons & Taylors used most commonly in water baths and in
ovens have a tolerance of plus or minus 3.6 degrees.
Procedure:
Verify lab or field thermometers against a calibrated thermometer. Check each thermometer
at two points if possible. Using water to perform the checks would be the best method.
It does no have to be a commercially made water bath. It could be as simple as using a
a bucket and adding hot water. A piece of Plexiglas with drilled holes works well when
placed over a water bath. Clothes pins or other clips can be used to hold the thermometers
in place while performing the check. Alternative methods may be used to achieve results.
Verification Intervals:
12 months
Report:
A copy must be retained for each thermometer tested and kept in the Testing Facilities
files for review.

## #64 Worksheet Thermometer
Thermometer
Procedure No. 64
ASTM E 77-07
Date:
Checked By:
Calibration Equipment Used:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks:
Standard Thermometer Serial No. |  |  |  |  | Laboratory Thermometer Serial No.
|  | ASTM No. |  |  |  |  | ASTM No.
| Standard Thermometer |  |  |  |  | Laboratory Thermometer
| Temperature Reading (°F or °C) |  |  |  |  | Temperature Reading (°F or °C)
| 1. |  |  |  |  | 1.
| 2. |  |  |  |  | 2.
| 3. |  |  |  |  | 3.
| 4. |  |  |  |  | 4.
Maximum scale allowed in accordance with ASTM E-1.
Remarks: